IMPROVEMENT OF A GAS PROVER UNCERTAINTY

Pier Giorgio SPAZZINI, Aline PICCATO, Paola PEDONE2

Abstract:

INRIM has operated the MICROGAS gas piston prover for small gas flow rates for several years now. This test rig allows measurements of flow rates lower than 0.1 SCCM with uncertainties of the order of 0.05% due to several technological features which include the two-sides temperature control of the gas chamber and the interferometric measurement of the piston displacement.

The most important uncertainty contribution in the overall uncertainty estimate for this test rig is the dimensional uncertainty of the piston (the complete uncertainty evaluation will be presented in the final paper).

After its initial realization, the MICROGAS piston was measured in the length section of the then-IMGC at three cross-sections. Due to the relatively low number of measured cross sections, it was decided to estimate the uncertainty of the diameter by considering the piston diameter as nominally constant with a conservative uncertainty value, derived from the uncertainties of the measurements.

At the beginning of 2013, the piston was measured again; improvements in the methodology at the length section of INRIM allowed to reduce the measurement uncertainty, but more interestingly, the measurement this time was performed at a larger number of cross sections. This allows to provide a more detailed description of the section evolution along the piston, and therefore a smaller uncertainty of the diameter at the various sections.

This improvement can be used to improve the overall uncertainty of the test rig by associating a different diameter, together with an associated uncertainty, to the various vertical positions measured for the piston.

The paper will describe the measurements performed on the piston diameter alongside with the method that will be used to describe the volume displacement evolution along the piston run and with its effect on the uncertainty estimate of the gas flow rate measurement.

The actual implementation of the method is planned for the near future.

# Footnotes

INRIM - Istituto Nazionale di Ricerca Metrologica, Strada delle Cacce 91, I-10135 Torino (ITALY) – Contact Author – Email p.spazzini@inrim.it

INRIM - Istituto Nazionale di Ricerca Metrologica, Strada delle Cacce 91, I-10135 Torino (ITALY)