## Slide 1: An Intercomparison of Water Flow and Gas Flow Laboratories using ISO 5167 Dp Devices
Jos van der Grinten NMi EuroLoop
Roy van Hartingsveldt NMi EuroLoop
Stefan Knöfel Siemens AG
Eberhard Gralla Siemens AG
Johann Kolar Seiko-Flowcontrol GmbH

## Slide 2: Contents
Introduction
Dp Packages
Results
  A and B tappings
  Cd versus Re
Discussion
Conclusions and recommendations

## Slide 3: Introduction
Combined Cycle Power Plants
Flow measurement equipment
  Venturis
  ISA 1932 nozzles
  Long radius nozzles
ISO 5167-3,4 (2003)
  Small range of Reynolds numbers

## Slide 4: Introduction – ISO characteristics
ReD /106 Orifice 0.005 ~ 100 (∞)
 | ISA 1932 Nozzle | Long radius Nozzle, high β | Venturi Machined
ISO 5167-.. | 3 | 3 | 4
β [-] | 0.44 ~ 0.80 | 0.25 ~ 0.80 | 0.40 ~ 0.75
d [mm] | 50 ~ 500 | 50 ~ 630 | 50 ~ 250
ReD /106 [-] | 0.02 ~ 10 | 0.01 ~ 10 | 0.2 ~ 1
Cd [-] | Eq. (3) | Eq. (8) | 0.995
U(Cd) [-] | β ≤ 0.6: 0.8% β >0.6: (2β-0.4)% | 2.0% | 1.0%

## Slide 5: Introduction – Questions
Questions
Scope expansion in Re domain
Improve uncertainty of the standard
Similarity of lab results
 Intercomparison
Fluid | Labs | t [ ̊C] | p [bara] | ReD [-]
Water | A,B,C.D | 20 / 40 | - | 7.2·105 - 2.5·106
Hot water | H | 80 | - | 2.5·106 - 1.0·107
Natural Gas | E,F,G | 10 - 20 | 45 - 50 | 3.5·106 - 2.4·107

## Slide 6: Dp packages
12” stainless steel, length 6 m
Nozzles β = 0.59, Venturi β = 0.64
ASME PTC-6 flow conditioner
Double pressure tappings A and B

## Slide 7: Results – 1
‑0.02% ± 0.09%

## Slide 8: Results – 2
‑0.04% ± 0.19%

## Slide 9: Results – 3
0.03% ± 0.21%

## Slide 10: Results – Calibration Cd – 1
Cd / Cd,ISO – 1 = – 0.21% ± 0.17%

## Slide 11: Results – Calibration Cd – 2
Cd / Cd,ISO – 1 = 0.09% ± 0.17%

## Slide 12: Results – Calibration Cd – 3
2%
3%

## Slide 13: Discussion – A & B tappings
Difference between A and B tappings
Generally small
Increase at the lower end of the Re range
ISA nozzle smallest differences Venturi greatest differences
Stronger effects at specific labs
  Trapped air in water pressure lines excluded
  Moisture in gas pressure lines excluded
  Other gas in gas pressure lines excluded
No influence on the general picture

## Slide 14: Discussion – Literature values
General
  All results within the uncertainty of the ISO 5167
  No discontinuities between water, hot water and natural gas
  Scope extension and uncertainty reduction possible
ISA 1932 nozzle
  Systematically below Cd,ISO
  Results may be approximate by a straight line
Long radius nozzle
  Water results close to literature values
  Scatter (2s value) identical to ISA
Venturi
  Highest scatter

## Slide 15: Discussion – Intercomparison
Intercomparison
Agreement if
95% confidence
ISA nozzle and long radius nozzle
  All lab results agree with at least 95% confidence
Venturi
  Gas results agree with two exceptions
  Water at the low Re some significant differences
  Inherent instability due to abrupt shape changes: stall
  Gaming technology? (Van Weers et al., Flomeko 1998)
  Design improvement (Reader-Harris et al., FM&I 2001)

## Slide 16: Conclusions – initial questions
Our questions:
Scope extension
  Possible for all devices
  Curve of water, hot water and natural gas connect very well
Uncertainty reduction
  Possible for all devices
  More β ratios to be tested
  For best uncertainty – calibrate in proper Re range
Intercomparison
  Agreement between labs for the nozzles
  Venturi: some significant differences
  ISA 1932 nozzle best suited for wide Re range

## Slide 17: Thank you !
EuroLoop: Calibration is our passion !
?