## Slide 1
NMISA
Bilateral Comparison between NMISA and VTT-MIKES - Part of the NMISA Flow Laboratory Journey to Accreditation
By : Deona Jonker
Richard Högström
Mpilo Dlamini
TQM-5126-1

## Slide 2: Introduction
NMISA is mandated to provide measurement traceability to the South African industry
NMISA gas flow laboratory performs the majority of gas flow calibrations – results in long calibration lead times
First priority for the NMISA gas flow laboratory – obtain ISO/IEC 17025 accreditation
  VTT-MIKES acted as reference laboratory for a bilateral comparison using four mass flow controllers as comparison artefacts
Second priority – obtain a second secondary reference standard
Using secondary standards as reference standards:
  Importing traceability from other overseas NMIs
  Standards must be calibrated at an overseas NMI
    Reference standard is absent from laboratory for ± two months

## Slide 3: The NMISA Gas Flow Laboratory
Part of the Flow Section in Physical Metrology Group
Gas flow (volume gas flow) calibration services are offered in the flow range 5,0 mL/min to 50 000 mL/min
Nitrogen gas is used as the calibration medium
Environmental conditions are monitored by a barometer and temperature- and humidity loggers
Typical instruments received for calibration include mass flow controllers, mass flow meters, bubble flow meters and rotameters

## Slide 4: The NMISA Gas Flow Laboratory
The gas flow laboratory is equipped with two secondary reference gas flow standards as highest-level standards
  Bios ML-800 positive displacement piston prover
  Fluke molbloc-L laminar flow elements (LFE)
The secondary reference standards are calibrated by overseas National Metrology Institutes with relevant calibration and measurement capabilities (CMCs) in the BIPM key comparison database (KCDB)
Traceability to international standards is therefore imported from these NMIs

## Slide 5: Road to Accreditation
First step
  the implementation of a quality system in the laboratory which adheres to the overall NMISA Total Quality Management System based on the requirements of the ISO/IEC 17025 standard
Second step
  A laboratory comparison had to be performed with an overseas National Metrology Institute (NMI) over the required flow range (5,0 mL/min to 50 000 mL/min) to support the laboratory’s proposed calibration and measurement capabilities (CMCs)
Third step
  Metrologists had to register with the NLA (National Laboratory Association) MetCert personnel certification scheme as metrologists or at the South African Council for Natural Scientific Professions (SACNSP) as scientists
    To be approved as a Technical signatory by SANAS

## Slide 6: Road to Accreditation
Laboratory underwent an initial assessment by SANAS on 22 & 23 March 2016
  An international expert was employed as technical assessor
  The laboratory received a positive recommendation for accreditation pending the successful clearance of the non-conformances identified

## Slide 7: MIKES – NMISA Bilateral Comparison
Comparison arrangements
  Four MKS mass flow controllers (50 mL/min, 500 mL/min, 5 000 mL/min, 50 000 mL/min) were used as comparison artefacts
  Measurements were first performed by NMISA
  The four mass flow controllers were then sent to the VTT-MIKES flow laboratory for calibration
  After VTT-MIKES had completed the calibration, the mass flow controllers were returned to NMISA for a second calibration
  VTT-MIKES acted as the reference laboratory (pilot), compiling the comparison report and generating the calibration certificates

## Slide 8: MIKES – NMISA Bilateral Comparison
Measurements were performed at the following nominal points:
  MFC(1): 5 mL/min; 12,5 mL/min; 25 mL/min; 37,5 mL/min and 50 mL/min
  MFC(2): 50 mL/min; 125 mL/min; 250 mL/min; 375 mL/min and 500 mL/min
  MFC(3): 500 mL/min; 1 250 mL/min; 2 500 mL/min; 3 750 mL/min and 5 000 mL/min
  MFC(4): 5 000 mL/min; 12 500 mL/min; 25 000 mL/min; 37 500 mL/min and 50 000 mL/min

## Slide 9: MIKES – NMISA Bilateral Comparison
Measurement results were reported at the standard conditions of 0 °C and 101,325 kPa
Nitrogen (purity of at least 99,999%) was used as the calibration gas
Mass flow controllers inlet pressure was set to 1,7 bar(g) (170 000 Pa)
Measurement uncertainty calculations were performed for each measurement
Specific excel spreadsheet was provided by the pilot laboratory for the reporting of the results
Calibration certificates were generated for the calibrations
A description and pictures of the calibration system had to be submitted to the pilot laboratory as well as the NMISA gas flow laboratory’s traceability information
NMISA submitted their two sets of measurement results simultaneously after the second calibration was completed
VTT-MIKES compiled the comparison report

## Slide 10: MIKES – NMISA Bilateral Comparison
Measurements performed at VTT-MIKES
  Calibrations below 30 000 mL/min were performed using the LFE (laminar flow element) calibration system
  Calibrations above 30 000 mL/min, the DWS2 primary standard was used
  LFE calibration system is calibrated against the DWS1 primary standard to establish traceability to SI units
  The two primary standards, DWS1 and DWS2, are primary gravimetric standards for gas mass flow which are traceable to SI units through measurements of mass and time.
  At each measurement point four one minute measurements were performed

## Slide 11: MIKES – NMISA Bilateral Comparison
Measurements performed at NMISA
  NMISA performed two sets of measurements, one before and one after VTT-MIKES measurements
  Mass flow controllers were calibrated against the Bios ML-800 reference gas flow standard
  Mass flow controllers were connected in line with the Bios
  A nitrogen gas cylinder was connected to the inlet of the flow bench with its outlet connected to the inlet of the mass flow controller
  The outlet of the mass flow controller supplied the reference standard with the flow medium
  Ten consecutive readings from the reference standard were acquired at each measuring point

## Slide 12: MIKES – NMISA Bilateral Comparison

## Slide 13: Method for Analysing the Results
At each measurement point, the relative calibration correction ( ) was determined as
The difference between the reference laboratory and the participating laboratory (DNMISA) was determined as

## Slide 14: Method for Analysing the Results
The uncertainty of the difference was calculated as
The drift udrift was determined as the largest difference between the weighted mean of the NMISA measurements and one single measurement assuming a rectangular probability distribution:

## Slide 15: Method for Analysing the Results
Significance of the observed differences
To analyse the significance of the differences calculated with equation (DNMISA), normalized error values (En) were calculated by dividing the difference with its expanded uncertainty (k=2)
If -1 < En < 1, the estimate for the difference is smaller than its expanded uncertainty
In this case, there is no statistically significant difference between the results obtained by the reference laboratory and the participating laboratory

## Slide 16: Method for Analysing the Results
Drift of the transfer standards
  The drift of the transfer standards during this comparison was studied by comparing the results of the measurement sets at NMISA to each other
  The drift was generally less than the measurement uncertainty for all transfer standards and therefore no corrections were made to the results
  Instead, the drift was included in the uncertainty of the NMISA results
  The differences in the uncertainties for the NMISA1 and NMISA2 measurements were due to poor reproducibility between the two sets of measurements

## Slide 17: Method for Analysing the Results
Chart: Nominal flow (mL/min)
| 5 | 12.5 | 25 | 37.5 | 50
NMISA1 (MFC1) | -6.5418 | -3.7522 | -4.1837 | -4.809 | -4.2579
NMISA1 (MFC2) | -10.0574 | -5.4133 | -3.4396 | -2.9354 | -2.8486
NMISA1 (MFC3) | 1.4343 | 0.18743 | 0.44748 | -0.11252 | 0.11585
NMISA1 (MFC4) | -16.9133 | -6.5071 | -3.2894 | -1.8388 | -1.074
NMISA2 (MFC1) | -11.5599 | -5.3432 | -4.8656 | -5.2599 | -4.5663
NMISA2 (MFC2) | -13.1778 | -6.0985 | -3.8001 | -3.3506 | -2.8954
NMISA2 (MFC3) | -0.9155 | -0.59285 | 0.064283 | -0.19966 | 0.26734
NMISA2 (MFC4) | -18.6386 | -6.8182 | -3.0888 | -1.8229 | -1.0352

## Slide 18: Comparison Results
 | MIKES |  | NMISA |  | MIKES – NMISA |  | 
Volume flow (mL/min) | (%) | (%) | (%) | (%) | (%) | (%) | En
MFC (1) |  |  |  |  |  |  | 
5 | -12.3 | 1.1 | -7.6 | 5.0 | -4.7 | 5.1 | -0.9
12.5 | -5.8 | 1.0 | -4.3 | 1.5 | -1.5 | 1.8 | -0.9
25 | -4.6 | 1.0 | -4.3 | 0.8 | -0.4 | 1.3 | -0.3
37.5 | -4.2 | 0.9 | -4.9 | 0.5 | 0.7 | 1.1 | 0.6
50 | -3.5 | 1.0 | -4.3 | 0.8 | 0.8 | 1.3 | 0.6
MFC(2) |  |  |  |  |  |  | 
50 | -14.5 | 2.1 | -10.2 | 4.1 | -4.3 | 4.6 | -0.9
125 | -7.2 | 1.0 | -5.5 | 1.1 | -1.7 | 1.5 | -1.2
250 | -4.5 | 1.0 | -3.5 | 0.6 | -1.0 | 1.1 | -0.9
375 | -3.8 | 1.0 | -3.0 | 0.5 | -0.8 | 1.1 | -0.7
500 | -3.1 | 0.7 | -2.9 | 0.5 | -0.2 | 0.9 | -0.3
MFC(3) |  |  |  |  |  |  | 
500 | -0.2 | 0.7 | 1.4 | 2.8 | -1.6 | 2.9 | -0.6
1250 | -0.5 | 0.8 | 0.2 | 0.9 | -0.7 | 1.2 | -0.6
2500 | -0.1 | 0.8 | 0.4 | 0.5 | -0.5 | 0.9 | -0.6
3750 | -0.4 | 0.8 | -0.1 | 0.2 | -0.3 | 0.8 | -0.4
5000 | -0.1 | 0.7 | 0.2 | 0.2 | -0.2 | 0.8 | -0.3
MFC(4) |  |  |  |  |  |  | 
5000 | -18.9 | 1.2 | -17.4 | 4.2 | -1.6 | 4.4 | -0.4
12500 | -8.2 | 0.7 | -6.6 | 0.7 | -1.6 | 1.0 | -1.6
25000 | -4.8 | 0.7 | -3.2 | 0.4 | -1.6 | 0.8 | -1.9
37500 | -3.4 | 1.1 | -1.8 | 0.4 | -1.6 | 1.1 | -1.4
50000 | -2.9 | 1.5 | -1.1 | 0.6 | -1.8 | 1.6 | -1.1

## Slide 19: Comparison Results
Chart: MFC(1)
| 5 | 12.5 | 25 | 37.5 | 50
MIKES | -12.2783 | -5.8463 | -4.6359 | -4.1767 | -3.5305
NMISA | -7.588 | -4.3239 | -4.2583 | -4.8692 | -4.3249

## Slide 20: Comparison Results
Chart: MFC(2)
| 50 | 125 | 250 | 375 | 500
MIKES | -14.5358 | -7.2277 | -4.4737 | -3.7798 | -3.1076
NMISA | -10.2303 | -5.4938 | -3.4945 | -3.0072 | -2.8707

## Slide 21: Comparison Results
Chart: MFC(3)
| 500 | 1250 | 2500 | 3750 | 5000
MIKES | -0.21482 | -0.54904 | -0.07381 | -0.42114 | -0.0502
NMISA | 1.3575 | 0.1761 | 0.43365 | -0.11747 | 0.17122

## Slide 22: Comparison Results
Chart: MFC(4)
| 5000 | 12500 | 25000 | 37500 | 50000
MIKES | -18.9296 | -8.1891 | -4.8403 | -3.4141 | -2.8843
NMISA | -17.3725 | -6.5687 | -3.2493 | -1.8308 | -1.0728

## Slide 23: Discussion of Comparison Results
The results of NMISA mainly agreed with the VTT-MIKES results for transfer standards MFC(1)-(3) in the volume flow range 5 mL/min to 5 000 mL/min
  All normalized error values were within -1 ≤ En ≤ 1, except for measurement point 125 mL/min with MFC(2)
  The observed discrepancy was probably due to the poor repeatability and stability of the MFCs when operated at low set point values
  This is supported by the fact that the difference between laboratories was smaller at 50 mL/min, 500 mL/min and 5000 mL/min for measurements performed with a MFC at set point of 100 % instead of 10 %

## Slide 24: Discussion of Comparison Results
A disagreement between the NMISA and VTT-MIKES results was found at flow rates from 12 500 mL/min to 50 000 mL/min, where En values were above one
  This discrepancy cannot be explained by the transfer standard, as it was found to perform well in this flow range
  A systematic difference of about -1,6 % was found in this flow range, which suggests that there might be an underestimated systematic error component affecting the results of NMISA

## Slide 25: Follow-up measurements to investigate discrepancies
Measurements were repeated for the 50 000 mL/min mass flow controller (MFC(4)) using the LFE calibration system as reference standard
The measurements were performed according to the comparison protocol
A nitrogen gas cylinder supplied gas through the reference LFE and the flow bench to the mass flow controller
The nitrogen gas was connected to the inlet of the LFE
The outlet of the LFE was connected to the inlet of the flow bench; the flow bench’s outlet was connected to the inlet of the mass flow controller
The outlet of the mass flow controller was open to atmosphere

## Slide 26: Follow-up measurements to investigate discrepancies
Chart: MFC(4)
| 5000 | 12500 | 25000 | 37500 | 50000
MIKES | -18.9296 | -8.1891 | -4.8403 | -3.4141 | -2.8843
NMISA1 | -16.9133 | -6.5071 | -3.2894 | -1.8388 | -1.074
NMISA2 | -18.6386 | -6.8182 | -3.0888 | -1.8229 | -1.0352
NMISA LFE | -18.1083 | -6.5122 | -3.1927 | -2.4896 | -2.2995

## Slide 27: Follow-up measurements to investigate discrepancies
Chart: MFC(4)
| 5000 | 12500 | 25000 | 37500 | 50000
MIKES-NMISA1 | -2.0163 | -1.682 | -1.5509 | -1.5753 | -1.8102
MIKES-NMISA2 | -0.29106 | -1.3709 | -1.7515 | -1.5912 | -1.8491
MIKES-NMISA_LFE | -0.82133 | -1.6769 | -1.6477 | -0.9245 | -0.58477

## Slide 28: Follow-up measurements to investigate discrepancies
The measurements performed with the LFE as reference standard were consistent with the results performed with the BIOS
  Based on these findings, it seems that the BIOS performance is not the reason for the discrepancy observed in the comparison
A possible reason for the observed discrepancies is the reproducibility and short-term stability of the transfer standard MFC(4)
  Due to the different reference standards used by NMISA and MIKES, the calibration cycles were different in terms of repetitions per calibration point and length of calibration cycle
  Also the laboratory temperature deviated by 4 °C
    This might have resulted in a different response of the mass flow controller

## Slide 29: Conclusion
To further investigate the discrepancies found at flow rates above 12 500 mL/min:
  Direct comparison measurements between the Bios ML-800 and Fluke laminar flow element reference standards should be performed and the results analysed
  Measurement uncertainties will be investigated
    Currently the main sources of uncertainty taken into consideration are the measurement uncertainty of the reference standard, repeatability and instrument resolution
    It seems not to be adequate enough - measurement uncertainty sources like reference standard stability, temperature and pressure influences will be investigated and included in the uncertainty budget
  Opportunities for NMISA to participate in further gas flow inter-laboratory comparisons should be investigated
    inter-comparisons using high accuracy transfer standards, such as piston provers or LFEs